## Slide 1
Detector Physics
Scintillation Detectors
D. Lascar | Postdoctoral Fellow | TRIUMF
GAPS Postdoc Lecture Series – November 16, 2015

## Slide 2: Outline
Scintillation Detectors
  Organic v. inorganic crystals
  Detection efficiency
  Light output response
Photomultiplier Tubes
  Photocathodes
  Electron optics
  Charge multiplication/dynodes

## Slide 3: From semiconductors to scintillators
Similarities:
  Takes advantage of the increased density of crystal lattices
  Makes use of band gap energies
  Doping lowers the energy required for a measureable interaction
Differences:
  Whereas scintillators re-radiate the energy absorbed as photons, semiconductors “move” electrons.

## Slide 4: Scintillators
Ionizing radiation ionizes and/or excites the matter it passes through.
When the excited matter returns to g.s. in some materials γ’s are in the visible range.
  Radioluminesence
The most efficient materials for visible γ generation are called scintillators.
If light emission continues for > 1 ms
  Phosphor
Notes: Phosphors are normally used as a thin layer coating the surface of a clear substrate.

## Slide 5: Energy Collection
Counters need only note that some energy was collected.
For calorimetery the goal is to convert the incident energy to a proportional amount of light.
  Losses from shower photons
  Losses from fluorescence x-rays
Slide courtesy of M. Fortner, N. Illinois University

## Slide 6: Photon Statistics
Typical Problem
Gamma rays at 450 keV are absorbed with 12% efficiency. Scintillator photons with average 2.8 eV produce photoelectrons 15% of the time.
What is the energy to produce a measurable photoelectron?
How does this compare to a gas detector (W-value)?
Answer
The total energy of scintillation is 450 x 0.12 = 54 keV.
  5.4 x 104 / 2.8 = 1.93 x 104 photons produced
  1.93 x 104 x 0.15 = 2900 photoelectrons produced
The equivalent W-value for the scintillator is:
  450 keV/2900 = 155 eV/pe
  W-value in gas = 30 eV/ip
Slide courtesy of M. Fortner, N. Illinois University

## Slide 7: Good Scintillator Requirements
High efficiency γ production.
Transparent to the λ of the emitted γ.
  n ~ 1.5
τ should decay quickly and with minimal delay.
Pulse shape discrimination
Cheap would be nice
Notes: Will go over pulse shape discrimination later

## Slide 8: Organic Scintillators
Made of linked or condensed benzene structures.
  Tend to be expensive.
Liquid scintillator made by dissolving material in solvent
Plastic scintillators radiate UV
  Cheaper
  Wavelength shifters (flour) required
https://en.wikipedia.org/wiki/Anthracene
Anthracene
Stilbene
https://en.wikipedia.org/wiki/(Z)-Stilbene
Notes: Anthracene and Stilbene are just two examples
In a plastic, the mean free path of the UV photon is only a few mm
Flour added at ~1% level to absorb UV photons and re-emit at longer wavelength (lower frequency)

## Slide 9: Pi-Bonds
Carbon in molecules has one excited electron.
  G.S. 1s22s22p2
  Excited 1s22s12p3
Hybrid p-orbitals are p-orbitals.
  Overlapping p-orbitals form bonds
  Appears in double bonds
Slide courtesy of M. Fortner, N. Illinois University

## Slide 10: Organic scintillator response
Above 125 keV e- response is linear
Proton/heavy particle always has a lower response
MeV electrons equivalent (MeVee)
  1 MeV e- → 1 MeVee
  >2 MeV p → 1 MeVee
Scintillator response (arb units)
Particle Energy (MeV)
NE-102
Notes: NE-102 is a plastic scintillator
Scintillator response differs by particle type and energy
Proton/heavy particle response begins an order of magnitude lower but that gap decreases

## Slide 11: Organic scintillator response
Flourescent energy/unit path length
Energy loss/unit path length
Normal scintillation efficiency
Notes: Physicist from Glascow, JB Birks suggested that high ionization density along particle track leads to quenching from damage molecules.
Damage leads to decreased efficiency.

## Slide 12: Organic scintillator response
To account for the probability of quenching:
Birk’s formula
Experimentally derived
Normalization parameter
Notes: If we assume that damage ionization density then density becomes
is a proportionality constant
Some fraction, , of the density becomes damaged
is experimentally derived by fitting to data
If we assume that damage ∝ ionization density then density becomes 𝐵(𝑑𝐸∕𝑑𝑥)
𝐵 is a proportionality constant
Some fraction, 𝑘, of the density becomes damaged
𝑘𝐵 is experimentally derived by fitting to data

## Slide 13: Organic scintillator response - electrons
To account for the probability of quenching:
For high-E e-, is small
The light output is linearly related to

## Slide 14: Organic scintillator response – α’s
To account for the probability of quenching:
For α’s,

## Slide 15: Organic Scintillators
Notes: Anthracene and Stilbene are just two examples
In a plastic, the mean free path of the UV photon is only a few mm
Flour added at ~1% level to absorb UV photons and re-emit at longer wavelength (lower frequency)

## Slide 16: Inorganic Scintillators
Fluorescence is known in many natural crystals.
  UV light absorbed
  Visible light emitted
Artificial scintillators can be made from many crystals.
  Doping impurities added
  Improve visible light emission
Higher density = More interaction = More efficient
Slide courtesy of M. Fortner, N. Illinois University

## Slide 17: Band Structure
Impurities in the crystal provide energy levels in the band gap.
Charged particles excites electrons to states below the conduction band.
De-excitation causes photon emission.
  Crystal is transparent at photon frequency.
conduction band
valence band
hn
impurity excited states
impurity ground state
Slide courtesy of M. Fortner, N. Illinois University

## Slide 18: Jablonski Diagram
Jablonski diagrams characterize the energy levels of the excited states.
  Vibrational transitions are low frequency
  Fluorescence and phosphorescence are visible and UV
Transitions are characterized by a peak wavelength lmax.
Slide courtesy of M. Fortner, N. Illinois University

## Slide 19: Time Lag
Fluorescence typically involves three steps.
  Excitation to higher energy state.
  E loss through change in vibrational state
  Emission of fluorescent photon.
The time for 1/e of the atoms to remain excited is the characteristic time t.
10-15 s
S1
S0
10-12 s
10-7 s
Slide courtesy of M. Fortner, N. Illinois University

## Slide 20: Decay Constant (τd)
Notes: N is the number of photons emitted in time, t.
N0 is the total number of photons emitted.

## Slide 21: Crystal Specs
www.detectors.saint-gobain.com
Common crystals are based on alkali halides
  Thallium or sodium impurities
Fluorite (CaF2) is a natural mineral scintillator.
(BGO, Bi4Ge3O12) is popular in physics detectors.
Crystal | τ (ns) | λmax (nm) | Output (nm)
Na(Tl) | 250 | 415 | 100
CsI(Tl) | 1000 | 550 | 45
CsI | 16 | 315 | 5
ZnS(Ag) | 130 | 110 | 450
CaF2(Eu) | 930 | 435 | 50
BGO | 300 | 480 | 20
Slide courtesy of M. Fortner, N. Illinois University
Notes: BGO is Bismuth germanate

## Slide 22: What can we do with photons
The radiation that interacts with the scintillator generates photons.
We can’t count photons.
We can manipulate them.
  Light guides can reflect and transmit photons at near 100% efficiency.
We can send them into materials to generate predictable behavior.
Photocathodes

## Slide 23: Photoelectron Emission
valence band
conduction band
Fermi energy
vacuum energy
EG
EA
y
hn
e-
Counting photons requires conversion to electrons.
The photoelectric effect can eject electrons from a material into a vacuum.
  Exceed gap energy and electron affinity energy
  Compare to work function
Slide courtesy of M. Fortner, N. Illinois University

## Slide 24: Quantum Efficiency
Reflection coefficient
Photon absorption
Mean escape length
Probability to eject from surface
Probability to reach vacuum energy
There is a probability that a photon will produce a free electron.
  Depends on bulk material properties & atomic properties
This is expressed as the quantum efficiency .
Slide courtesy of M. Fortner, N. Illinois University
Notes: Photocathodes are designed to maximize the quantum efficiency.
  Layer of semiconductor or alkali compound on glass.
Quantum efficiency dominated by L and Ps.
  Thin material that passes electrons easily for L
  Material with low EA to improve Ps

## Slide 25: Commercial Photocathodes
Different photocathodes vary in response to and .
  Alkali for UV detection (Cs-I, Cs-Te)
  Bialkali for visible light (Sb-Rb-Cs, Sb-K-Cs)
  Semiconductors for visible to IR (GaAsP, InGaAs)
Hamamatsu.com
Slide courtesy of M. Fortner, N. Illinois University

## Slide 26: Challenge:
Single electrons are hard to detect. What to do?
Stronger Field
  Gives electrons more energy
  If the electrons have enough energy to ionize along their path
Constant E-field so avalanche everywhere
ΔV
i
PPAC
late
arallel
valanche
ounter
Notes: Remember this from the last lecture? It’s the same problem. Single electrons are hard to detect. What to do?

## Slide 27: Electron Multiplier
Electrons can be multiplied by interaction with a surface.
  Emitter: BeO, GaP
  Metal substrate: Ni, Fe, Cu
This electrode is called a dynode.
substrate electrode
e
emissive surface

## Slide 28: Photomultiplier Tube (PMT)
PMTs combines a photocathode and series of dynodes.
High voltage is divided between the dynodes.
Output current is measured at the anode.
  Sometimes at the last dynode
Slide courtesy of M. Fortner, N. Illinois University

## Slide 29: Gain ()
depends on the material and .
  typically 0.7-0.8
Multiple dynodes are staged to increase gain.
  Photocathode current
  Input stage current
Total gain is a product of stage gain.
  Collection efficiency
Slide courtesy of M. Fortner, N. Illinois University

## Slide 30: Amplifier
Photomultiplier tubes often have 10-14 stages.
  Gain in excess of 107
A single photon can produce a measurable charge.
  Single photoelectron
  ~ 10-12 C
Fast response in about 1 ns.
  ~ 1 mA
Slide courtesy of M. Fortner, N. Illinois University

## Slide 31: Dark Current
Phototubes have “dark” current even with no incident light.
  Thermionic emission
  Anode leakage
  Case scintillation
  Gas ionization
This increases with applied voltage.
Slide courtesy of M. Fortner, N. Illinois University

## Slide 33: Flux
Radiant flux (W) measures the energy flowing at a point per unit time.
Luminous flux (lm) weights the flux for its impact on a visual system.
  Peak efficiency constant Km = 683 lm/W
  Spectral luminous efficiency V(l)

## Slide 34: Luminous Area
The flux per unit area is the irradiance (illuminance) E.
  lm/m2 = lux
Flux from a luminous area is the emittance M.
  Lux not used
dA
dA
dF
dF
Emittance
Irradiance

## Slide 35: Intensity
Intensity relates to the flux from a point source.
  Flux per unit solid angle
  Definition of the candela
Intensity is calibrated by a current from a known source with a known response s.

## Slide 36: Noise
Dark current contributes to the noise in a measurement.
  Equivalent noise input
  For Df = 1 Hz, P ~ 10-15 W
Signal to noise depends on the statistical fluctuations, dark current and readout circuit.
  Dominated by statistics
Slide courtesy of M. Fortner, N. Illinois University

## Slide 37: Position Response
Photocathodes are not uniform.
  Variations in response on the surface.
Angular response falls off beyond 40°.